## Slide 1: Sheep 11 Round 1
07/25/2023
RR | MHR (bpm) | MSpO2 (proximal) (%) | FHR (bpm) | FSaO2 (%)
15 | 88 | 94 | 119 | 37.3

## Slide 2: Optical Characteristics
Lock-in frequencies For LEDs:
  740nm – 930Hz​ (50%)
  850nm – 690Hz (50%)
Drive current for LEDs:
  740nm – 400mA ​(213.33mW peak & 106.67mW average optical power)
  850nm – 300mA​ (420mW peak & 210mW average optical power)
 | CH1 | CH2 | CH3 | CH4 | CH5
Detector distance | 1.5 cm | 3 cm | 4.5 cm | 7 cm | 10 cm
Detector Area | 0.051 cm2 | 0.051 cm2 | 0.051 cm2 | 1 cm2 | 1 cm2
Detector Responsivity

## Slide 3: Electrical Characteristics
Down sampled to 80hz
IQ demodulated from 690Hz & 930Hz
 | CH1 | CH2 | CH3 | CH4 | CH5
TIA Gain | 10k | 100k | 1M | 10M | 100M
PGA Gain | 1x | 2x | 2x | 1x | 2x

## Slide 4
Round 1 starting 01:05:16 PM extracted 0-5 minutes Det 1&5
MHR
MHR
MHR
MHR
MHR
FHR
Fetal signal in Ch5 (850nm)
Two FHR signals?